## 輸出摘要
此文件從 Numbers 輸出。已將每個表格轉換為不同的 Excel 工作表。每個 Numbers 工作表上的所有其他物件將放置在不同的工作表上。請注意，公式計算可能與 Excel 不同。
Numbers 工作表名稱 | Numbers 表格名稱 | Excel 工作表名稱
Daily rev0 (+CN)
| 表格 1 | Daily rev0 (+CN)
Daily rev0
| 表格 1 | Daily rev0

## Daily rev0 (+CN)
Startup | [6 empty] | Rev.00
Place | Sub system | Device | Check items | (Unit) | Standard value | Check | Remark
Accelerator Control room | Beam Scheduler | Network status | PLC & PC network status | ---- | ON_LINE | □ON_LINE / □OFF_LINE
|  | Cyrostat | Maximum Upper coil tempareture | (K) | < 6.7 K | / / / | Upper coil : A / B / C/ D
|  |  | Maximum power lead tempareture | (K) | < 65 K |  | Power lead: P / N
|  | Cyclotron vacuum chamber | Vacuum pressure(before startup)(A) | (Pa) | <2.0E-4
|  |  | (After pass water)(B) The difference pressure between(B) and (A) | (Pa) | <0.5E-5
|  | Sector | Temperature | (℃) | < 33 ℃
|  | Filament | Filament current/ operation times | (A)/(Mins) | >90A/<2000mins
|  | Chimney | Operation time (arc ON) | (Mins)
|  | Cooling water | Water flow rate | ---- |  | □Ready / □ Not ready
|  | Cooling water | Purity | ---- |  | □No alarm / □ Alarm
|  | Cryopump-A | 2nd temperature / 1st temperature | (K) | < 20 / < 85
|  | Cryopump-B | 2nd temperature / 1st temperature | (K) | < 20 / < 85
Accelerator room & pit | Accelerator | Cryopump | Water flow rate | (L/min) | >9
|  | Cryocooler | Water flow rate (Comp A, B, C, D) | (L/min) | >9 | / / /
|  | Gas cylinder (H2) | Hydrogen gas cylinder valve | ---- | ---- | □Open the valve
|  | Ion sourece | Exchange the polarity of filament current | ---- | ---- | □Exchange the cable | should be done Every morning
| Building, Accelerator & BTS | Building | Accelerator room temperature/Hunidity | (℃)/(%) | 22-26℃/ < 60% | /
|  | Building | BT ambient temperature/Humidity | (℃)/(%) | 22-26℃/ < 60% | /
|  | Cooling water manifold | Water leakage | ---- | No leakage | □No leakage / □ leakage
|  | Door | No one remains in accelerator room and close the shilding door | ---- | No one | □No one & close the door
Cooling water room | Accelerator | Gas cylinder (Dry air) | Dry air gas cylinder valve | ---- | ---- | □Open the valve
| Building | Pumps | Noise | ---- | No abnomal noise | □No abnomal noize / □abnomal noize
Power supply room | Building Power supply | Building | Power supply room temperature/humidity | (℃)/(%) | 22-26℃/ < 60% | /
|  |  | Water leakage Check | ---- | No leakage | □No leakage / □ Leakage
|  | Distribution panel | Power-receiving switch | ---- | ON | □On / □Off
|  |  | Cable | ---- | No damage | □No damage / □Damage
Daily Startup Sheet
| Startup | Supervisor: Operator: .
Date: / /
Time: :
shutdown | [6 empty] | Rev.02
Place | Sub system | Device | Check items | (Unit) | Standard value | Check | Remark
Accelerator Control room | Beam Scheduler terminal | Integrated beam use interlock screen | Integrated beam use Ion chamber | ---- | ---- | □Record
|  | Cyclotron vacuum chamber | Vacuum pressure with gas | (Pa) | < 2.0E-3
|  | Cyrostat | Maximum Upper coil tempareture | (K) | < 6.7 K |  | Upper coil : A / B / C/ D
|  |  | Maximum power lead tempareture | (K) | < 65 K |  | Power lead: P / N
|  | Cryopump-A | 2nd temperature / 1st temperature | (K) | < 20 / < 85
|  | Cryopump-B | 2nd temperature / 1st temperature | (K) | < 20 / < 85
Cooling water room | Accelerator | Gas cylinder | Dry air gas cylinder | ---- | ---- | □Close the valve
| Building | Pumps | Noise | ---- | No abnomal noise | □No abnomal noize / □abnomal noize
Accelerator room & pit | Accelerator Building | Gas cylinder | Hydrogen gas cylinde | ---- | ---- | □Close the valve
|  | Building | Accelerator room temperature/Hunidity | (℃)/(%) | 22-26℃/ < 60% | /
|  |  | BT ambient temperature/Humidity | (℃)/(%) | 22-26℃/ < 60% | /
|  |  | Water leakage Check | ---- | No leakage | □No leakage / □ Leakage
Power supply room | Building Power supply | Building | Power supply room temperature/humidity | (℃)/(%) | 22-26℃/ < 60% | /
|  |  | Water leakage Check | ---- | No leakage | □No leakage / □ Leakage
Daily Shotdown Sheet
| Shutdown | Supervisor: Date: / / Time: : Operator: .

## Daily rev0
Startup | [6 empty] | Rev.00
Place | Sub system | Device | Check items | (Unit) | Standard value | Check | Remark
Accelerator Control room | Beam Scheduler | Network status | PLC & PC network status | ---- | ON_LINE | □ON_LINE / □OFF_LINE
|  | Cyrostat | Maximum Upper coil tempareture | (K) | < 6.7 K |  | Upper coil : A / B / C/ D
|  |  | Maximum power lead tempareture | (K) | < 65 K |  | Power lead: P / N
|  | Cyclotron vacuum chamber | Vacuum pressure(before startup)(A) | (Pa) | <2.0E-4
|  |  | (After pass water)(B) The difference pressure between(B) and (A) | (Pa) | <0.5E-5
|  | Sector | Temperature | (℃) | < 33 ℃
|  | Filament | Filament current/ operation times | (A)/(Mins) | >90A/<2000mins
|  | Chimney | Operation time (arc ON) | (Mins)
|  | Cooling water | Water flow rate | ---- |  | □Ready / □ Not ready
|  | Cooling water | Purity | ---- |  | □No alarm / □ Alarm
|  | Cryopump-A | 2nd temperature / 1st temperature | (K) | < 20 / < 85
|  | Cryopump-B | 2nd temperature / 1st temperature | (K) | < 20 / < 85
Accelerator room & pit | Accelerator | Cryopump | Water flow rate | (L/min) | >9
|  | Cryocooler | Water flow rate (Comp A, B, C, D) | (L/min) | >9 | / / /
|  | Gas cylinder (H2) | Hydrogen gas cylinder valve | ---- | ---- | □Open the valve
|  | Ion sourece | Exchange the polarity of filament current | ---- | ---- | □Exchange the cable | should be done Every morning
| Building, Accelerator & BTS | Building | Accelerator room temperature/Hunidity | (℃)/(%) | 22-26℃/ < 60% | /
|  | Building | BT ambient temperature/Humidity | (℃)/(%) | 22-26℃/ < 60% | /
|  | Cooling water manifold | Water leakage | ---- | No leakage | □No leakage / □ leakage
|  | Door | No one remains in accelerator room and close the shilding door | ---- | No one | □No one & close the door
Cooling water room | Accelerator | Gas cylinder (Dry air) | Dry air gas cylinder valve | ---- | ---- | □Open the valve
| Building | Pumps | Noise | ---- | No abnomal noise | □No abnomal noize / □abnomal noize
Power supply room | Building Power supply | Building | Power supply room temperature/humidity | (℃)/(%) | 22-26℃/ < 60% | /
|  |  | Water leakage Check | ---- | No leakage | □No leakage / □ Leakage
|  | Distribution panel | Power-receiving switch | ---- | ON | □On / □Off
|  |  | Cable | ---- | No damage | □No damage / □Damage
Daily Startup Sheet
| Startup | Supervisor: Date: / / Time: : Operator: .
shutdown | [6 empty] | Rev.02
Place | Sub system | Device | Check items | (Unit) | Standard value | Check | Remark
Accelerator Control room | Beam Scheduler terminal | Integrated beam use interlock screen | Integrated beam use Ion chamber | ---- | ---- | □Record
|  | Cyclotron vacuum chamber | Vacuum pressure with gas | (Pa) | < 2.0E-3
|  | Cyrostat | Maximum Upper coil tempareture | (K) | < 6.7 K |  | Upper coil : A / B / C/ D
|  |  | Maximum power lead tempareture | (K) | < 65 K |  | Power lead: P / N
|  | Cryopump-A | 2nd temperature / 1st temperature | (K) | < 20 / < 85
|  | Cryopump-B | 2nd temperature / 1st temperature | (K) | < 20 / < 85
Cooling water room | Accelerator | Gas cylinder | Dry air gas cylinder | ---- | ---- | □Close the valve
| Building | Pumps | Noise | ---- | No abnomal noise | □No abnomal noize / □abnomal noize
Accelerator room & pit | Accelerator Building | Gas cylinder | Hydrogen gas cylinde | ---- | ---- | □Close the valve
|  | Building | Accelerator room temperature/Hunidity | (℃)/(%) | 22-26℃/ < 60% | /
|  |  | BT ambient temperature/Humidity | (℃)/(%) | 22-26℃/ < 60% | /
|  |  | Water leakage Check | ---- | No leakage | □No leakage / □ Leakage
Power supply room | Building Power supply | Building | Power supply room temperature/humidity | (℃)/(%) | 22-26℃/ < 60% | /
|  |  | Water leakage Check | ---- | No leakage | □No leakage / □ Leakage
Daily Shotdown Sheet
| Shutdown | Supervisor: Date: / / Time: : Operator: .